PART V — INSTRUMENT TRAYS CARDIOVASCULAR TRAYS Complete, illustrated examination notes General · Gynaecologic & Obstetric · Genitourinary · Thoracic Cardiovascular · Orthopaedic · Neurologic · ENT · Ophthalmic · Paediatric Legal (8.5 × 14 in) · 1.27 cm margins · monochrome print-ready Every figure is an original line diagram, drawn for clarity in black and white

Table of Contents

To fill in the page numbers: click anywhere in the list below, press Ctrl+A then F9, and choose “Update entire table”.

Right-click here and choose “Update Field”, or press Ctrl+A then F9, to build the Table of Contents with page numbers.

PART V # Cardiovascular Trays Vascular procedures · Vascular shunt · Cardiac procedures

Cardiovascular trays obey rules found nowhere else in surgery. Every instrument that touches a vessel must be atraumatic, because a crushed intima thromboses. Every suture is non-absorbable monofilament on a swaged, usually double-armed needle. Heparinised saline is on the tray from the start. And the cardiac tray adds the single greatest complexity in surgery — cardiopulmonary bypass.

## 5.1   Vascular Procedures Tray

Also called:  General vascular tray · arterial reconstruction set

| The vascular tray serves any procedure in which an artery or vein is opened, repaired, bypassed or replaced. Its design follows one principle: the vessel wall must be controlled without being injured. Every clamp is non-crushing, every forceps has atraumatic jaws, and the tray carries the means to occlude, open, flush, repair and graft. Used for / includes ▪ Aorto-iliac and femoro-popliteal bypass grafting ▪ Carotid endarterectomy ▪ Abdominal aortic aneurysm repair (open) ▪ Embolectomy and thrombectomy ▪ Repair of arterial trauma ▪ Arteriovenous fistula and graft creation for dialysis ▪ Varicose vein surgery (the venous subset) |  | Fig 5.1 — Everything is atraumatic Atraumatic means what? ▪ DeBakey jaws have fine longitudinal rows — they hold without crushing the intima. ▪ A toothed forceps on an artery is a surgical error: intimal injury seeds thrombosis. ▪ The same logic makes every vascular clamp non-crushing.. |
| --- | --- | --- |

### Contents of the tray

| Qty | Instrument | Notes / rationale |
| --- | --- | --- |
| BASE SET |  |  |
| 1 | Major procedures tray | for exposure of the vessel |
| 1 | Long instruments tray | for aortic and iliac work |
| VASCULAR CLAMPS — ALL NON-CRUSHING |  |  |
| 6 | Bulldog clamps, assorted (DeBakey, Dietrich) | signature item — small spring clips for temporary occlusion of small vessels |
| 4 | Satinsky (side-biting) clamps | signature item — allows partial occlusion so flow continues while a patch or graft is sewn on |
| 4 | Angled DeBakey aortic clamps | cross-clamping the aorta |
| 4 | Fogarty / Cooley atraumatic clamps | soft, insert-jawed clamps |
| 2 | Carotid / Javid clamps | carotid surgery |
| — | Vessel loops (silastic) and umbilical tapes | encircle and gently occlude |
| — | Rubber-shod (shodded) clamps | protects the vessel and graft |
| — | Bulldog clamp applier / holder |  |
| ATRAUMATIC GRASPING AND DISSECTING |  |  |
| 6 | DeBakey vascular forceps, 15/20/24/30 cm | signature item — longitudinal atraumatic rows, not teeth |
| 2 | Gerald forceps, fine | delicate handling |
| 2 | Potts-Smith angled vascular scissors, 45° and 60° | signature item — extends an arteriotomy accurately |
| 2 | Fine curved Metzenbaum / dissecting scissors |  |
| 2 | Freer / Penfield elevators | endarterectomy plane dissection |
| 1 | Endarterectomy spatula / dissector | raises the atheromatous plaque |
| 6 | Long right-angled (Mixter) forceps | encircling vessels |
| 2 | Nerve hook / vein hook | lifting side branches |
| OPENING, FLUSHING AND GRAFTING |  |  |
| 1 | Knife handle No.7 with No.11 blade | the arteriotomy is started with a No.11 |
| 1 | Aortic punch, sized set | signature item — cuts a clean circular hole for a graft anastomosis |
| — | Heparinised saline with bulb syringe and blunt irrigating cannulae | signature requirement — flushes and prevents thrombus at every stage |
| 2 | Fogarty embolectomy (balloon) catheters, sized | signature item — withdraws thrombus and embolus |
| — | Graft material — PTFE, Dacron, or reversed autologous vein | sized to the vessel |
| 1 | Vein stripper and vein harvesting set | long saphenous vein harvest |
| 1 | Valvulotome | in-situ vein bypass |
| 1 | Tunneller | creates the subcutaneous graft tunnel |
| — | Graft sizers and calibrators |  |
| — | Topical haemostatic agents and fibrin sealant |  |
| SUTURING — A DISTINCT DISCIPLINE |  |  |
| 4 | Fine needle holders (Castroviejo, DeBakey pattern) | delicate, often spring-handled |
| — | Non-absorbable monofilament (polypropylene) suture 3-0 to 8-0, double-armed | signature requirement — never absorbable, never braided, in an artery |
| — | Teflon pledgets | buttress sutures in a friable or calcified wall |
| 2 | Fine suture scissors |  |
| — | Bulldog clamps on suture ends | keeps the suture ends controlled |
| 1 | Magnifying loupes | small-vessel anastomosis |
| MONITORING, IMAGING AND PHARMACOLOGY |  |  |
| 1 | Doppler probe (sterile) and machine | confirms flow before closure |
| 1 | Image intensifier with contrast | completion angiography |
| — | Systemic heparin and protamine | anticoagulation and its reversal |
| — | Papaverine / vasodilator | relieves spasm |
| — | Cross-matched blood and a cell saver |  |
| — | Pressure monitoring lines |  |
| — | Closed suction drain |  |

#### Additional supplies and accessories

Position — Depends on the vessel — supine for aortic, carotid (neck extended, turned away) and femoral; prone or lateral rarely.

Golden rules — Proximal control before distal. Obtain control of the vessel above the lesion before opening it. Never grasp the intima. Flush before final closure to expel air and debris.

Complications — Haemorrhage, thrombosis of the repair or graft, distal embolisation, graft infection, anastomotic pseudoaneurysm, reperfusion injury and compartment syndrome.

Examination pointers ▪ Open the answer with the principle: all instruments atraumatic, all sutures non-absorbable monofilament, heparinised saline throughout. ▪ Name DeBakey forceps, Satinsky clamp, bulldog clamp, Potts-Smith scissors, Fogarty catheter, aortic punch — the six vascular signatures. ▪ Satinsky = partial (side-biting) occlusion, so distal flow is maintained — a favourite distinction. ▪ Fogarty catheter for embolectomy and heparin/protamine are reliable marks. ▪ State proximal control before distal and never grasp the intima. ▪ Teflon pledgets for a friable wall is a detail that distinguishes a strong answer.

Common mistakes to avoid ✗ Suggesting absorbable or braided suture for an arterial anastomosis. ✗ Using toothed forceps on a vessel. ✗ Forgetting heparinised saline and systemic heparin. ✗ Omitting completion assessment (Doppler or angiography).

#### Likely examination questions

Q1.  Enumerate the contents of a vascular procedures tray.

A.  A major and long instruments tray plus: non-crushing clamps — bulldog, Satinsky side-biting, angled DeBakey aortic, Fogarty/Cooley, Javid carotid clamps, vessel loops and rubber-shod clamps; atraumatic DeBakey forceps in several lengths, Gerald forceps, Potts-Smith angled scissors, endarterectomy spatula and elevators, Mixter forceps and vein hooks; No.11 blade, sized aortic punch, heparinised saline with bulb syringe and cannulae, Fogarty embolectomy catheters, PTFE/Dacron/vein graft material, vein stripper, valvulotome and tunneller; fine Castroviejo needle holders with double-armed polypropylene 3-0 to 8-0 and Teflon pledgets; plus sterile Doppler, image intensifier with contrast, heparin and protamine, cross-matched blood, cell saver and a drain.

Q2.  Why is a Satinsky clamp particularly useful?

A.  It is a side-biting clamp that occludes only part of the vessel circumference, so a patch or graft can be sewn to the wall while blood continues to flow through the remaining lumen — avoiding complete occlusion and distal ischaemia.

## 5.2   Vascular Shunt Tray

Also called:  Portosystemic / arteriovenous shunt tray · dialysis access set

| A shunt deliberately diverts blood from one vessel to another. Two families exist: portosystemic shunts for portal hypertension (portocaval, splenorenal, mesocaval), and arteriovenous shunts/fistulae for haemodialysis access. Both are built on the vascular tray, with additions for the specific vessels and shunt material. Used for / includes ▪ Portocaval, mesocaval and distal splenorenal (Warren) shunt for variceal bleeding ▪ Transjugular intrahepatic portosystemic shunt (TIPS) — endovascular equivalent ▪ Radiocephalic (Brescia-Cimino) and brachiocephalic arteriovenous fistula ▪ Arteriovenous graft using PTFE for dialysis access ▪ Temporary intraluminal shunt in vascular trauma to restore flow ▪ Carotid shunt during endarterectomy |  | Fig 5.2 — Two vessels, one anastomosis Selective versus non-selective ▪ Non-selective (portocaval, mesocaval): decompresses the whole portal system → effective but high encephalopathy rate. ▪ Selective (distal splenorenal / Warren): decompresses only the gastro-oesophageal varices → less encephalopathy. ▪ A very commonly examined distinction.. |
| --- | --- | --- |

### Contents of the tray

| Qty | Instrument | Notes / rationale |
| --- | --- | --- |
| BASE SET |  |  |
| 1 | Complete vascular procedures tray | see 5.1 — the entire set is required |
| 1 | Major and long instruments trays | for portosystemic (abdominal) shunts |
| SHUNT MATERIAL AND SIZING |  |  |
| — | PTFE / Dacron graft, 6–10 mm | interposition portosystemic shunt or AV graft |
| — | Reversed autologous vein | the preferred conduit where suitable |
| 2 | Temporary intraluminal (Javid, Pruitt-Inahara) shunts | signature item — maintains distal perfusion while the anastomosis is constructed or during trauma repair |
| — | Graft sizers, calibrators and tunneller |  |
| — | Shunt clamps and shunt-securing tapes |  |
| 1 | Balloon shunt with inflation syringe | carotid shunting |
| PORTOSYSTEMIC SHUNT ADDITIONS |  |  |
| 2 | Satinsky clamps, large | partial occlusion of the inferior vena cava |
| 4 | Angled DeBakey / Cooley clamps | portal vein and superior mesenteric vein |
| 2 | Deep Deaver and malleable retractors | retroperitoneal exposure |
| 1 | Self-retaining ring retractor |  |
| — | Portal pressure manometry set | measures the pressure gradient before and after shunting |
| 2 | Long right-angled forceps | encircling the portal vein |
| — | Vessel loops, wide | portal vein and IVC |
| 1 | Head-light |  |
| — | Splenectomy instruments available | for a splenorenal shunt |
| ARTERIOVENOUS ACCESS ADDITIONS |  |  |
| 1 | Fine (microvascular) instrument set | radial artery and cephalic vein are 2–4 mm |
| — | Magnifying loupes or an operating microscope | essential for small-vessel anastomosis |
| — | Fine polypropylene 6-0 to 8-0, double-armed |  |
| 2 | Fine bulldog clamps and micro-clamps |  |
| 1 | Vessel dilators, fine, graded | gently dilates a spastic vessel |
| — | Papaverine / lignocaine topical | relieves vasospasm |
| 1 | Sterile Doppler probe | confirms a thrill/bruit before closure |
| — | Tourniquet and arm-board |  |
| — | Heparinised saline with fine cannulae |  |
| ASSESSMENT AND CLOSURE |  |  |
| 1 | Sterile Doppler / flow probe | patency and flow direction confirmed |
| 1 | Image intensifier with contrast | completion angiography or shuntogram |
| — | Pressure monitoring lines and manometer |  |
| — | Heparin and protamine |  |
| — | Cross-matched blood and cell saver | portal hypertension patients bleed and are coagulopathic |
| — | Closed suction drains |  |
| — | Fresh frozen plasma, platelets, vitamin K | coagulopathy of liver disease |

#### Additional supplies and accessories

Position — Supine for portocaval (right subcostal or midline) and for AV access (arm on an arm-board, abducted). Left flank/thoracoabdominal for splenorenal shunts.

Physiological caution — A portosystemic shunt diverts portal blood past the liver: it relieves variceal bleeding but precipitates hepatic encephalopathy. Selective shunts (distal splenorenal) decompress varices while preserving portal perfusion and cause less encephalopathy.

Maturation of an AV fistula — A new fistula needs about 6 weeks to mature before it can be cannulated; a graft can be used sooner.

Examination pointers ▪ Divide the answer into portosystemic and arteriovenous shunts — the two families have different additions. ▪ The temporary intraluminal shunt (Javid, Pruitt-Inahara) is the signature item; explain that it maintains distal perfusion. ▪ For portosystemic shunts include a portal pressure manometry set. ▪ For AV access include loupes/microscope and fine 6-0 to 8-0 suture — the vessels are only millimetres wide. ▪ State the selective vs non-selective distinction and hepatic encephalopathy as the key complication. ▪ Doppler confirmation of flow before closure applies to both families.

#### Likely examination questions

Q1.  What is a vascular shunt tray and how does it differ according to the type of shunt?

A.  It is a complete vascular tray plus shunt-specific additions. For portosystemic shunts: large Satinsky and angled DeBakey/Cooley clamps for the IVC and portal vein, PTFE or Dacron interposition graft, deep retractors and a ring retractor, long Mixter forceps, wide vessel loops, a portal pressure manometry set, head-light and blood products for coagulopathy. For arteriovenous access: a fine microvascular set with loupes or a microscope, fine bulldog and micro-clamps, graded vessel dilators, topical papaverine, 6-0 to 8-0 double-armed polypropylene, heparinised saline and a sterile Doppler. Both need temporary intraluminal shunts, heparin and protamine and completion imaging.

Q2.  Why does a non-selective portosystemic shunt cause encephalopathy?

A.  It diverts the entire portal venous flow — carrying nitrogenous products absorbed from the gut — directly into the systemic circulation, bypassing hepatic detoxification. Selective shunts decompress only the varices and preserve portal perfusion of the liver, so encephalopathy is less frequent.

## 5.3   Cardiac Procedures Tray

Also called:  Open-heart tray · CABG and valve replacement set · cardiopulmonary bypass tray

| The cardiac tray is the largest and most complex set in surgery. It combines a sternotomy set, a complete vascular set, the cannulae and lines for cardiopulmonary bypass, instruments for the specific intracardiac procedure, and everything needed to restart and support the heart. Its organising idea is that the circulation is handed over to a machine, the heart stopped, repaired, and restarted. Used for / includes ▪ Coronary artery bypass grafting (CABG) ▪ Aortic and mitral valve replacement or repair ▪ Repair of congenital defects — ASD, VSD, tetralogy of Fallot ▪ Ascending aortic and aortic arch replacement; aortic dissection ▪ Cardiac tumour excision; pulmonary thromboendarterectomy ▪ Heart transplantation and mechanical support device implantation |  | Fig 5.3 — Open the sternum, take over the circulation, repair, restart Count the wire ▪ Sternal wires, needle fragments and valve sizers are all countable items. ▪ A broken wire fragment in the mediastinum is a classic retained surgical item. ▪ An emergency re-sternotomy set stays at the bedside in intensive care for tamponade.. |
| --- | --- | --- |

### Contents of the tray

| Qty | Instrument | Notes / rationale |
| --- | --- | --- |
| BASE SETS |  |  |
| 1 | Complete major procedures tray |  |
| 1 | Complete vascular procedures tray | see 5.1 |
| 1 | Long instruments tray |  |
| STERNOTOMY — OPENING AND CLOSING THE CHEST |  |  |
| 1 | Sternal saw (oscillating or reciprocating) with spare blades | signature item — median sternotomy |
| 1 | Lebsche knife and mallet | manual alternative if the saw fails |
| 1 | Sternal (Finochietto / Ankeney) self-retaining retractor | spreads the divided sternum |
| 1 | Internal mammary artery (IMA) retractor | elevates the sternal edge to harvest the IMA |
| — | Bone wax | sternal marrow haemostasis |
| — | Sternal wires (No.5/No.6) with a heavy needle | sternal closure |
| 1 | Wire twister, wire cutter and wire holder | must be counted — wire fragments are retained items |
| 1 | Sternal punch / awl | wire passage |
| 1 | Sternal plating set | for a high-risk sternum |
| CARDIOPULMONARY BYPASS — THE DEFINING GROUP |  |  |
| 1 | CPB circuit: oxygenator, reservoir, heat exchanger, arterial pump and tubing | primed and checked by the perfusionist before the skin incision |
| 2 | Aortic cannula with introducer | arterial return to the aorta |
| 2 | Venous cannulae — single two-stage or bicaval | drains the right atrium / cavae |
| 1 | Antegrade cardioplegia cannula | into the aortic root |
| 1 | Retrograde cardioplegia cannula | into the coronary sinus |
| 1 | Aortic root vent / needle vent | removes air from the aortic root |
| 1 | Left atrial / ventricular vent | decompresses the left heart |
| 4 | Aortic cross-clamp and caval (Satinsky) clamps | cross-clamping isolates the heart |
| — | Purse-string sutures with tourniquets (snuggers) | secure each cannula at its insertion site |
| — | Cardioplegia solution, cold, with delivery set | arrests the heart in diastole and protects the myocardium |
| — | Topical cold saline / ice slush | surface cooling |
| — | Systemic heparin and protamine | full heparinisation before cannulation; reversal after bypass |
| — | ACT (activated clotting time) monitoring |  |
| 1 | Cardiotomy suckers (two) with tubing | returns shed blood to the circuit |
| CORONARY BYPASS (CABG) ADDITIONS |  |  |
| 1 | Internal mammary artery harvesting set | fine dissecting scissors, clips, IMA retractor |
| 1 | Long saphenous vein harvesting set | open or endoscopic; vein stripper, fine clips |
| 1 | Radial artery harvest set |  |
| 2 | Coronary (Beaver / micro) knife and blades | arteriotomy in a 1.5–2 mm coronary artery |
| 2 | Potts-Smith coronary scissors, 45° / 60° | extending the arteriotomy |
| 1 | Coronary dilators and probes, graded | sizes the vessel |
| 1 | Aortic punch, sized | proximal anastomosis on the aorta |
| 2 | Coronary bulldog / occluders and silastic slings | controls the coronary during anastomosis |
| 1 | CO₂ / saline blower-mister | clears the anastomotic field |
| — | Polypropylene 7-0 / 8-0 double-armed | coronary anastomosis |
| — | Magnifying loupes | essential |
| VALVE PROCEDURE ADDITIONS |  |  |
| 1 | Valve retractors and leaflet retractors | exposes the annulus |
| 1 | Valve leaflet excision scissors and rongeur | excises the diseased valve |
| 1 | Annulus debridement / calcium rongeur |  |
| 1 | Valve sizers / obturators, complete set | signature item — selects the prosthesis size |
| 1 | Valve holder and rotator with handle |  |
| — | Prosthetic valves — mechanical and bioprosthetic, full size range | confirmed present before bypass begins |
| — | Annuloplasty rings and sizers | mitral valve repair |
| — | Pledgeted mattress sutures, 2-0 | annular fixation |
| 1 | Nerve hook and valve tester |  |
| RESTARTING AND SUPPORTING THE HEART |  |  |
| 1 | Internal (sterile) defibrillator paddles | signature requirement — defibrillates the heart directly |
| — | Temporary epicardial pacing wires and external pacemaker | atrial and ventricular |
| 1 | De-airing needle and vent | expels intracardiac air |
| 1 | Transoesophageal echocardiography probe | assesses repair and ventricular function |
| — | Inotropes and vasoactive infusions, prepared |  |
| — | Intra-aortic balloon pump available | low cardiac output |
| — | Ventricular assist device available | for failure to wean from bypass |
| 2 | Mediastinal and pleural chest drains with underwater seal | always left in place |
| — | Cell saver, cross-matched blood, platelets, FFP, cryoprecipitate |  |
| — | Emergency re-sternotomy set kept at the bedside | for post-operative tamponade — a critical requirement |

#### Additional supplies and accessories

Position — Supine, arms tucked; legs prepared and exposed for vein harvest; wide skin preparation from chin to ankles.

Team — Surgeon, assistants, anaesthetist, perfusionist, scrub and circulating nurses. The perfusionist primes and checks the bypass circuit before the incision — an integral part of tray preparation.

Sequence — Sternotomy → harvest conduits → heparinise → cannulate aorta and right atrium → go on bypass → cross-clamp aorta → deliver cardioplegia → perform the repair → de-air → release the cross-clamp → rewarm → defibrillate/pace as needed → wean from bypass → decannulate → protamine → haemostasis → drains → sternal wires.

Complications — Bleeding and re-operation, cardiac tamponade, low cardiac output, arrhythmia, stroke and neurocognitive injury, renal failure, sternal dehiscence and mediastinitis, and heparin-induced thrombocytopenia.

Mnemonic — the five systems of a cardiac tray ▪ "S-B-P-V-R" ▪ Sternotomy (saw, retractor, wax, wires) ▪ Bypass (cannulae, circuit, cardioplegia, clamps) ▪ Procedure — CABG (conduits, punch, coronary set) or valve (sizers, prostheses) ▪ Vascular set (DeBakey, Satinsky, Potts, fine suture) ▪ Restart (internal paddles, pacing wires, TOE, inotropes, drains)

Examination pointers ▪ Structure the answer as five systems — sternotomy, bypass, procedure-specific, vascular, restart. This is the only way to cover it completely. ▪ Cardiopulmonary bypass items are the defining group: aortic and venous cannulae, cardioplegia cannulae, aortic cross-clamp, purse-string sutures with tourniquets, cardioplegia solution, heparin/protamine. ▪ Cardioplegia arrests the heart in diastole and protects the myocardium — state the purpose, not just the name. ▪ Internal defibrillator paddles and temporary epicardial pacing wires for restarting the heart are commonly forgotten. ▪ Sternal wires with wire twister and cutter — and the fact that they must be counted. ▪ Mention the perfusionist and that the circuit is primed before incision. ▪ Emergency re-sternotomy set at the bedside post-operatively is a high-value detail.

Common mistakes to avoid ✗ Listing instruments without the bypass circuit and cannulae. ✗ Forgetting how the heart is restarted (internal paddles, pacing wires). ✗ Omitting sternal closure items — wax, wires, twister. ✗ Not confirming prosthesis availability before bypass starts.

#### Likely examination questions

Q1.  Describe the preparation of a cardiac procedures tray.

A.  It comprises five systems. Sternotomy: sternal saw with spare blades, Lebsche knife, sternal self-retaining and IMA retractors, bone wax, sternal wires with twister, cutter and punch. Cardiopulmonary bypass: a primed circuit with oxygenator, reservoir and heat exchanger; aortic and venous cannulae; antegrade and retrograde cardioplegia cannulae; root and ventricular vents; aortic cross-clamp and caval clamps; purse-string sutures with tourniquets; cold cardioplegia; heparin, protamine and ACT monitoring; cardiotomy suckers. Procedure-specific: for CABG — IMA, saphenous and radial harvest sets, coronary knife, Potts scissors, coronary dilators, aortic punch, blower-mister, 7-0/8-0 polypropylene and loupes; for valves — valve retractors, leaflet scissors, calcium rongeur, a complete set of sizers, valve holder, prostheses and annuloplasty rings with pledgeted sutures. Vascular set in full. Restart and support: internal defibrillator paddles, temporary epicardial pacing wires with external pacemaker, de-airing needle, TOE, inotropes, IABP availability, mediastinal and pleural drains, cell saver and blood products, and an emergency re-sternotomy set for the bedside.

Q2.  What is the purpose of cardioplegia and how is it delivered?

A.  Cardioplegia is a cold, high-potassium solution that arrests the heart in diastole, abolishing electromechanical activity and greatly reducing myocardial oxygen demand so the myocardium tolerates the ischaemic period of cross-clamping. It is delivered antegradely into the aortic root above the cross-clamp and/or retrogradely through a cannula in the coronary sinus, and is supplemented by topical cooling.